Analysis of sulfur in VG-2 glass

The following analyses demonstrate the effect of trace sulfur analyses using different primary standards. So long as the sulfur standard was peaked for each standard (to account for peak shift) the analysis of the VG-2 glass was essentially consistent.

The only problem observed was that the composition for the Pyrrhotite standard was entered nominally as FeS. Once the actual analysis was entered into the standard database, this standard gave consistent analyses with the other standards.

Conclusion: It doesn’t matter what standard one uses for sulfur analyses provided that the analyst is measuring the peak intensity AT THE ACTUAL PEAK POSITIONS for both the unknown and standard.

VG-2 glass S analysis (each standard peaked individually)- NOTE, this measurement is not intended to be accurate, only as a basis for understanding the change in composition as a function of peak position and standard.

Standard	 S (weight percent, analyses at 15 kev, 10 sec)

Anhydrite: 	.14452

Pyrite		.14470

Pyrrhotite	.13285 (assumed 50/50)

Pyrrhotite	.14968 (Fe(1-x)S, where x=0.17) (from web database)

Pyrrhotite	.14472 (Fe(1-x)S, where x=0.13) (from EPMA analysis of actual standard)

St  327 Set   2 Anhydrite (CaSO4) UC # 5555

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

"slightly soluble in alcohol", JJD

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 05:24:01 PM to 11/11/2004 05:35:03 PM

Average Total Oxygen:       47.047     Average Total Weight%:  100.467

Average Calculated Oxygen:  47.047     Average Atomic Number:   13.533

Average Excess Oxygen:        .000     Average Atomic Weight:   22.765

Average ZAF Iteration:        2.40     Average Quant Iterate:     4.00

Results in Elemental Weight Percents

SPEC:        O      Sr

TYPE:     CALC    SPEC

AVER:  47.0466  .49300

SDEV:   .17662  .00000

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -3.46  -15.05  -24.16  -52.20     .00   -1.55   -6.06     .00     .00     .00     .00  -36.62

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

25 29.6721  .00081  .00393  .00000  .00694  .03721  .02486 23.4164  .00226  .00000  .00000  .01000 100.700

26 29.4445  .00053  .00000  .00000  .01163  .00079  .06265 23.5105  .02034  .00000  .05065  .00868 100.735

27 29.2846  .00000  .00715  .01199  .00000  .00000  .00000 23.6730  .04068  .01783  .00000  .00833 100.829

28 29.2147  .00125  .00501  .00197  .00000  .00000  .02409 23.4141  .00226  .00000  .00000  .00091 99.9872

29 29.0576  .00741  .00000  .00000  .00000  .00000  .02966 23.4902  .04746  .01019  .00000  .00000 100.083

AVER:  29.3347  .00200  .00322  .00279  .00371  .00760  .02825 23.5008  .02260  .00560  .01013  .00558 100.467

SDEV:   .23433  .00306  .00316  .00521  .00535  .01656  .02243  .10548  .02108  .00813  .02265  .00473

SERR:   .10480  .00137  .00141  .00233  .00239  .00741  .01003  .04717  .00943  .00364  .01013  .00212

%RSD:       .8   152.8    98.1   186.7   144.0   217.9    79.4      .4    93.3   145.2   223.6    84.8

PUBL:  29.1270   n.a.    n.a.    n.a.    n.a.    n.a.    n.a.  23.4990   n.a.    n.a.    n.a.    n.a.  100.000

%VAR:      .71     .00     .00     .00     .00     .00     .00     .01     .00     .00     .00     .00

DIFF:   .20770  .00000  .00000  .00000  .00000  .00000  .00000  .00184  .00000  .00000  .00000  .00000

STDS:      358     298     374     336     374     395      22     327     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .2205   .1599   .7502   .7341   .4736

STCT:   2807.0  6645.5  2659.3  1728.1  1650.2  3309.5  1241.8  1625.0   778.3  3452.2  3325.8 16139.1

UNKF:    .2694   .0000   .0000   .0000   .0000   .0001   .0002   .2205   .0002   .0000   .0001   .0000

UNCT:   4466.3      .5      .5     -.2     -.1     -.1      .5  1625.0     1.0     -.5     -.1     1.1

UNBG:     13.3    29.7    22.4     8.4    11.5     6.3     1.8     9.2     2.5     9.5     4.6    15.0

ZCOR:   1.0888  1.1725  1.3609  2.1509  1.0560  1.1978  1.2326  1.0659  1.1090  1.1810  1.2226  1.5829

KRAW:   1.5911   .0001   .0002  -.0001  -.0001   .0000   .0004  1.0000   .0013  -.0002   .0000   .0001

PKBG:   337.05    1.02    1.02     .97     .99     .99    1.25  179.04    1.40     .96    1.00    1.08

INT%:      .00     .00     .00     .00     .00    -.97     .00     .00     .00     .00     .00     .00

St  730 Set   1 Pyrite UC # 21334

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

Leadville, Colorado

Excellent standard for Fe and S

AA (by C. Lewis): Cu=0.005%, Au=0.00%, Ag=0.00%, Zn=0.004%

XRF (by J. Hampel): Ni=0.00%, Co=0.00%, Cu=0.01%, As=0.009%

Zn=0.0009%, Pb=0.010%, Ga=0.0001%, Cr=0.00%, Ti=0.058%

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 05:38:02 PM to 11/11/2004 05:49:06 PM

Average Total Oxygen:         .000     Average Total Weight%:  100.869

Average Calculated Oxygen:    .000     Average Atomic Number:   20.700

Average Excess Oxygen:        .000     Average Atomic Weight:   40.089

Average ZAF Iteration:        3.00     Average Quant Iterate:     4.00

Results in Elemental Weight Percents

SPEC:        O

TYPE:     SPEC

AVER:   .00000

SDEV:   .00000

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -9.87  -23.94  -34.49  -62.14     .00   -1.14   -5.03     .00     .00     .00     .00  -47.62

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

30  .00699  .00000  .00000  .01618  .01439 47.4054  .00000 53.6074  .00000  .00000  .05036  .00673 101.107

31  .00000  .00000  .00000  .00000  .00000 47.6958  .02756 53.7226  .00000  .03799  .00000  .00419 101.488

32  .00000  .00000  .00000  .00000  .00000 46.9164  .01803 53.1672  .00000  .00000  .00000  .00000 100.102

33  .00575  .00000  .00000  .00446  .00000 47.1962  .00000 53.5088  .00000  .00000  .01385  .00000 100.729

34  .00345  .00000  .00000  .00000  .01817 47.7225  .00000 53.1702  .00000  .00000  .00000  .00373 100.918

AVER:   .00324  .00000  .00000  .00413  .00651 47.3872  .00912 53.4352  .00000  .00760  .01284  .00293 100.869

SDEV:   .00322  .00000  .00000  .00701  .00902  .34135  .01293  .25484  .00000  .01699  .02181  .00291

SERR:   .00144  .00000  .00000  .00314  .00403  .15266  .00578  .11397  .00000  .00760  .00975  .00130

%RSD:     99.4      .1      .1   169.8   138.5      .7   141.8      .5      .1   223.6   169.9    99.3

PUBL:    n.a.    n.a.    n.a.    n.a.    n.a.  46.5500  .05800 53.4500   n.a.   .00000   n.a.    n.a.  100.058

%VAR:      .00     .00     .00     .00     .00    1.80  -84.28    -.03     .00     .00     .00     .00

DIFF:   .00000  .00000  .00000  .00000  .00000  .83723 -.04888 -.01477  .00000  .00000  .00000  .00000

STDS:      358     298     374     336     374     395      22     327     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .2205   .1599   .7502   .7341   .4736

STCT:   2807.0  6643.3  2659.3  1728.1  1650.2  3309.5  1241.8  1625.0   778.3  3452.2  3325.8 16139.1

UNKF:    .0000   .0000   .0000   .0000   .0001   .4358   .0001   .5057   .0000   .0001   .0001   .0000

UNCT:       .0    -1.9    -1.4     -.1      .0  2127.4     -.2  3727.1     -.9     -.5     -.3      .0

UNBG:     20.1    42.4    29.6    11.4    16.2    11.2     3.0    16.4     3.9    14.8     7.7    19.7

ZCOR:   1.0623  1.2091  1.4587  2.5127  1.1374  1.0874  1.0816  1.0566  1.0952  1.1014  1.1112  1.7741

KRAW:    .0000  -.0003  -.0005  -.0001   .0000   .6428  -.0002  2.2936  -.0011  -.0001  -.0001   .0000

PKBG:     1.00     .96     .95     .99    1.01  191.00     .93  228.26     .77     .97     .98    1.00

INT%:      .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00

St  757 Set   1 FeS (Pyrrhotite)

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 05:52:05 PM to 11/11/2004 06:03:10 PM

Average Total Oxygen:         .000     Average Total Weight%:  100.241

Average Calculated Oxygen:    .000     Average Atomic Number:   22.064

Average Excess Oxygen:        .000     Average Atomic Weight:   43.201

Average ZAF Iteration:        2.20     Average Quant Iterate:     4.00

Results in Elemental Weight Percents

SPEC:        O

TYPE:     SPEC

AVER:   .00000

SDEV:   .00000

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -8.54  -28.35  -39.62  -66.87     .00    -.86   -4.23     .00     .00     .00     .00  -52.94

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

35  .00000  .00000  .00000  .00000  .00000 60.0981  .01260 39.8057  .00468  .79053  .03079  .00456 100.747

36  .01076  .00000  .00000  .01006  .00000 59.5700  .00798 39.3829  .00234  .72710  .01908  .00626 99.7365

37  .00000  .00000  .00000  .00424  .00000 60.2258  .00000 39.4886  .00000  .85375  .04937  .00000 100.622

38  .00493  .00000  .00492  .00000  .00000 59.1870  .05884 39.5832  .01403  .80724  .00327  .00337 99.6668

39  .00396  .00000  .00192  .00000  .00000 59.9621  .00334 39.6697  .00468  .78788  .00000  .00000 100.434

AVER:   .00393  .00000  .00137  .00286  .00000 59.8086  .01655 39.5860  .00514  .79330  .02050  .00284 100.241

SDEV:   .00443  .00000  .00215  .00443  .00000  .42567  .02411  .16280  .00533  .04546  .02036  .00279

SERR:   .00198  .00000  .00096  .00198  .00000  .19036  .01078  .07281  .00238  .02033  .00911  .00125

%RSD:    112.8      .1   157.4   154.7      .0      .7   145.7      .4   103.6     5.7    99.3    98.2

PUBL:    n.a.    n.a.    n.a.    n.a.    n.a.  59.1110   n.a.  40.8890   n.a.    n.a.    n.a.    n.a.  100.000

%VAR:      .00     .00     .00     .00     .00    1.18     .00   -3.19     .00     .00     .00     .00

DIFF:   .00000  .00000  .00000  .00000  .00000  .69759  .00000 -1.3030  .00000  .00000  .00000  .00000

STDS:      358     298     374     336     374     395      22     327     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .2205   .1599   .7502   .7341   .4736

STCT:   2807.0  6641.1  2659.3  1728.1  1650.2  3309.5  1241.8  1624.9   778.3  3452.2  3325.8 16139.1

UNKF:    .0000   .0000   .0000   .0000   .0000   .5622   .0002   .3686   .0000   .0073   .0002   .0000

UNCT:       .3    -4.1    -1.5     -.4    -2.7  2744.4      .3  2716.6      .2    33.6      .8      .2

UNBG:     22.4    44.4    30.4    10.9    18.9    12.2     3.1    16.3     2.9    15.2     7.7    19.0

ZCOR:   1.0253  1.2751  1.5660  2.8313  1.0956  1.0639  1.0402  1.0739  1.1482  1.0880  1.0873  1.9502

KRAW:    .0001  -.0006  -.0006  -.0002  -.0016   .8293   .0002  1.6718   .0002   .0097   .0002   .0000

PKBG:     1.01     .91     .95     .97     .86  226.24    1.09  168.98    1.05    3.21    1.11    1.01

INT%:      .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00

St  298 Set   2 Glass, VG-2 USNM 111240/52

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

Juan de Fuca Basaltic Glass

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 04:55:55 PM to 11/11/2004 05:06:58 PM

Average Total Oxygen:       43.790     Average Total Weight%:   99.343

Average Calculated Oxygen:  43.790     Average Atomic Number:   12.809

Average Excess Oxygen:        .000     Average Atomic Weight:   22.070

Average ZAF Iteration:        3.00     Average Quant Iterate:     3.00

Results in Elemental Weight Percents

SPEC:        O

TYPE:     CALC

AVER:  43.7897

SDEV:   .10341

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -3.48  -21.61  -26.61  -49.62     .00    -.29   -2.50     .00     .00     .00     .00  -35.87

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

15 7.65472 23.7959 7.39518 1.98211  .16231 9.12237  .96959  .16737  .08598  .00507  .04794 3.99781 99.1249

16 7.64928 23.7442 7.39269 1.96991  .17020 8.81813 1.13551  .14267  .09453  .02029  .13018 4.00919 98.9828

17 7.67020 23.6588 7.42492 1.93467  .15238 9.02185 1.09244  .13833  .08592  .00000  .21747 4.07332 99.1742

18 7.70077 23.7526 7.39470 2.10572  .15194 9.14024 1.02480  .14393  .09455  .04818  .14112 4.10122 99.6728

19 7.77203 23.7885 7.42427 2.08671  .17793 9.05015 1.03496  .13028  .12601  .01014  .20417 4.02589 99.7580

AVER:  7.68940 23.7480 7.40635 2.01582  .16295 9.03055 1.05146  .14452  .09740  .01674  .14818 4.04149 99.3425

SDEV:   .05034  .05459  .01668  .07572  .01130  .12849  .06409  .01385  .01656  .01911  .06775  .04409

SERR:   .02251  .02441  .00746  .03386  .00505  .05746  .02866  .00619  .00741  .00855  .03030  .01972

%RSD:       .7      .2      .2     3.8     6.9     1.4     6.1     9.6    17.0   114.2    45.7     1.1

PUBL:  7.94739 23.7501 7.44119 1.94365  .15773 9.20329 1.10907  .13500  .08728   n.a.   .17038 4.04631 99.7554

%VAR:    -3.25    -.01    -.47    3.71    3.31   -1.88   -5.19    7.05   11.59     .00  -13.03    -.12

DIFF:  -.25799 -.00213 -.03484  .07217  .00522 -.17274 -.05761  .00952  .01012  .00000 -.02220 -.00482

STDS:      358     298     374     336     374     395      22     327     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .2205   .1599   .7502   .7341   .4736

STCT:   2807.0  6650.6  2659.3  1728.1  1650.2  3309.5  1241.8  1625.7   778.3  3452.2  3325.8 16139.1

UNKF:    .0713   .1848   .0532   .0100   .0015   .0766   .0090   .0012   .0007   .0001   .0012   .0261

UNCT:   1181.8  6650.0  2017.9   233.9    21.5   373.8    20.1     8.6     3.4      .6     5.6   888.5

UNBG:     12.7    26.6    21.3     8.5     9.8     5.9     1.6     3.3     1.7     8.2     4.4    14.3

ZCOR:   1.0787  1.2848  1.3920  2.0257  1.1032  1.1792  1.1716  1.2418  1.4051  1.1759  1.1996  1.5501

KRAW:    .4210   .9999   .7588   .1354   .0131   .1130   .0162   .0053   .0043   .0002   .0017   .0551

PKBG:    93.74  250.75   95.75   28.67    3.21   64.45   13.24    3.68    3.09    1.08    2.35   63.00

INT%:      .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00

St  298 Set   2 Glass, VG-2 USNM 111240/52

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

Juan de Fuca Basaltic Glass

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 04:55:55 PM to 11/11/2004 05:06:58 PM

Average Total Oxygen:       43.790     Average Total Weight%:   99.343

Average Calculated Oxygen:  43.790     Average Atomic Number:   12.809

Average Excess Oxygen:        .000     Average Atomic Weight:   22.070

Average ZAF Iteration:        3.00     Average Quant Iterate:     3.00

Results in Elemental Weight Percents

SPEC:        O

TYPE:     CALC

AVER:  43.7900

SDEV:   .10339

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -3.48  -21.61  -26.61  -49.62     .00    -.29   -2.50     .00     .00     .00     .00  -35.87

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

15 7.65472 23.7958 7.39518 1.98211  .16231 9.12237  .96959  .16760  .08598  .00507  .04794 3.99781 99.1254

16 7.64928 23.7442 7.39268 1.96991  .17020 8.81813 1.13551  .14286  .09453  .02029  .13018 4.00919 98.9832

17 7.67020 23.6588 7.42491 1.93467  .15238 9.02186 1.09244  .13850  .08592  .00000  .21748 4.07332 99.1746

18 7.70077 23.7526 7.39470 2.10572  .15194 9.14024 1.02480  .14411  .09455  .04818  .14112 4.10122 99.6733

19 7.77203 23.7885 7.42427 2.08671  .17793 9.05015 1.03496  .13043  .12601  .01014  .20417 4.02589 99.7584

AVER:  7.68940 23.7480 7.40635 2.01582  .16295 9.03055 1.05146  .14470  .09740  .01674  .14818 4.04149 99.3430

SDEV:   .05034  .05459  .01668  .07572  .01130  .12849  .06409  .01387  .01656  .01911  .06775  .04409

SERR:   .02251  .02441  .00746  .03386  .00505  .05746  .02866  .00620  .00741  .00855  .03030  .01972

%RSD:       .7      .2      .2     3.8     6.9     1.4     6.1     9.6    17.0   114.2    45.7     1.1

PUBL:  7.94739 23.7501 7.44119 1.94365  .15773 9.20329 1.10907  .13500  .08728   n.a.   .17038 4.04631 99.7554

%VAR:    -3.25    -.01    -.47    3.71    3.31   -1.88   -5.19    7.18   11.59     .00  -13.03    -.12

DIFF:  -.25799 -.00214 -.03484  .07217  .00522 -.17274 -.05761  .00970  .01012  .00000 -.02220 -.00482

STDS:      358     298     374     336     374     395      22     730     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .5061   .1599   .7502   .7341   .4736

STCT:   2807.0  6650.6  2659.3  1728.1  1650.2  3309.5  1241.8  3727.1   778.3  3452.2  3325.8 16139.1

UNKF:    .0713   .1848   .0532   .0100   .0015   .0766   .0090   .0012   .0007   .0001   .0012   .0261

UNCT:   1181.8  6650.0  2017.9   233.9    21.5   373.8    20.1     8.6     3.4      .6     5.6   888.5

UNBG:     12.7    26.6    21.3     8.5     9.8     5.9     1.6     3.3     1.7     8.2     4.4    14.3

ZCOR:   1.0787  1.2848  1.3920  2.0257  1.1032  1.1792  1.1716  1.2418  1.4051  1.1759  1.1996  1.5501

KRAW:    .4210   .9999   .7588   .1354   .0131   .1130   .0162   .0023   .0043   .0002   .0017   .0551

PKBG:    93.74  250.75   95.75   28.67    3.21   64.45   13.24    3.68    3.09    1.08    2.35   63.00

INT%:      .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00

St  298 Set   2 Glass, VG-2 USNM 111240/52

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

Juan de Fuca Basaltic Glass

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 04:55:55 PM to 11/11/2004 05:06:58 PM

Average Total Oxygen:       43.773     Average Total Weight%:   99.315

Average Calculated Oxygen:  43.773     Average Atomic Number:   12.810

Average Excess Oxygen:        .000     Average Atomic Weight:   22.071

Average ZAF Iteration:        3.00     Average Quant Iterate:     3.00

Results in Elemental Weight Percents

SPEC:        O

TYPE:     CALC

AVER:  43.7732

SDEV:   .10410

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -3.48  -21.61  -26.61  -49.62     .00    -.29   -2.50     .00     .00     .00     .00  -35.87

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

15 7.65461 23.7966 7.39517 1.98201  .16231 9.12229  .97014  .15387  .08599  .00507  .04794 3.99784 99.0933

16 7.64919 23.7448 7.39265 1.96981  .17019 8.81807 1.13605  .13116  .09454  .02029  .13018 4.00922 98.9559

17 7.67011 23.6594 7.42488 1.93457  .15238 9.02179 1.09299  .12716  .08592  .00000  .21747 4.07335 99.1482

18 7.70068 23.7532 7.39468 2.10562  .15194 9.14016 1.02534  .13231  .09455  .04818  .14112 4.10125 99.6457

19 7.77194 23.7891 7.42423 2.08661  .17793 9.05008 1.03549  .11975  .12602  .01014  .20417 4.02592 99.7335

AVER:  7.68931 23.7486 7.40632 2.01572  .16295 9.03048 1.05200  .13285  .09740  .01674  .14818 4.04151 99.3153

SDEV:   .05035  .05461  .01667  .07572  .01130  .12849  .06409  .01274  .01656  .01911  .06775  .04409

SERR:   .02252  .02442  .00746  .03386  .00505  .05746  .02866  .00570  .00741  .00855  .03030  .01972

%RSD:       .7      .2      .2     3.8     6.9     1.4     6.1     9.6    17.0   114.2    45.7     1.1

PUBL:  7.94739 23.7501 7.44119 1.94365  .15773 9.20329 1.10907  .13500  .08728   n.a.   .17038 4.04631 99.7554

%VAR:    -3.25    -.01    -.47    3.71    3.31   -1.88   -5.15   -1.59   11.60     .00  -13.03    -.12

DIFF:  -.25808 -.00150 -.03487  .07207  .00522 -.17281 -.05707 -.00215  .01012  .00000 -.02220 -.00480

STDS:      358     298     374     336     374     395      22     757     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .3387   .1599   .7502   .7341   .4736

STCT:   2807.0  6650.5  2659.3  1728.1  1650.2  3309.5  1241.9  2716.6   778.3  3452.2  3325.8 16139.1

UNKF:    .0713   .1848   .0532   .0100   .0015   .0766   .0090   .0011   .0007   .0001   .0012   .0261

UNCT:   1181.8  6649.9  2017.9   233.9    21.5   373.8    20.1     8.6     3.4      .6     5.6   888.5

UNBG:     12.7    26.7    21.4     8.5     9.8     5.9     1.6     3.3     1.7     8.2     4.4    14.3

ZCOR:   1.0786  1.2848  1.3920  2.0257  1.1032  1.1792  1.1716  1.2418  1.4052  1.1759  1.1996  1.5501

KRAW:    .4210   .9999   .7588   .1353   .0131   .1130   .0162   .0032   .0043   .0002   .0017   .0551

PKBG:    93.78  250.16   95.51   28.64    3.21   64.45   13.33    3.68    3.09    1.08    2.35   63.01

INT%:      .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00

St  298 Set   2 Glass, VG-2 USNM 111240/52

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

Juan de Fuca Basaltic Glass

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 04:55:55 PM to 11/11/2004 05:06:58 PM

Average Total Oxygen:       43.797     Average Total Weight%:   99.355

Average Calculated Oxygen:  43.797     Average Atomic Number:   12.809

Average Excess Oxygen:        .000     Average Atomic Weight:   22.070

Average ZAF Iteration:        3.00     Average Quant Iterate:     3.00

Results in Elemental Weight Percents

SPEC:        O

TYPE:     CALC

AVER:  43.7971

SDEV:   .10309

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -3.48  -21.61  -26.61  -49.62     .00    -.29   -2.50     .00     .00     .00     .00  -35.87

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

15 7.65478 23.7955 7.39514 1.98212  .16231 9.12240  .96960  .17337  .08598  .00507  .04794 3.99781 99.1392

16 7.64933 23.7439 7.39265 1.96991  .17020 8.81816 1.13552  .14778  .09453  .02029  .13018 4.00919 98.9950

17 7.67025 23.6585 7.42488 1.93468  .15238 9.02188 1.09245  .14327  .08592  .00000  .21748 4.07332 99.1860

18 7.70082 23.7523 7.39466 2.10572  .15194 9.14026 1.02480  .14907  .09454  .04818  .14112 4.10122 99.6851

19 7.77207 23.7882 7.42423 2.08672  .17794 9.05017 1.03496  .13492  .12601  .01014  .20417 4.02589 99.7691

AVER:  7.68945 23.7477 7.40631 2.01583  .16295 9.03058 1.05147  .14968  .09740  .01674  .14818 4.04148 99.3549

SDEV:   .05034  .05458  .01668  .07572  .01130  .12850  .06409  .01435  .01656  .01911  .06775  .04409

SERR:   .02251  .02441  .00746  .03386  .00505  .05746  .02866  .00642  .00741  .00855  .03030  .01972

%RSD:       .7      .2      .2     3.8     6.9     1.4     6.1     9.6    17.0   114.2    45.7     1.1

PUBL:  7.94739 23.7501 7.44119 1.94365  .15773 9.20329 1.10907  .13500  .08728   n.a.   .17038 4.04631 99.7554

%VAR:    -3.25    -.01    -.47    3.71    3.31   -1.88   -5.19   10.88   11.59     .00  -13.03    -.12

DIFF:  -.25794 -.00241 -.03488  .07218  .00522 -.17271 -.05761  .01468  .01011  .00000 -.02220 -.00483

STDS:      358     298     374     336     374     395      22     757     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .3816   .1599   .7502   .7341   .4736

STCT:   2807.0  6650.6  2659.3  1728.1  1650.2  3309.5  1241.8  2716.6   778.3  3452.2  3325.8 16139.1

UNKF:    .0713   .1848   .0532   .0100   .0015   .0766   .0090   .0012   .0007   .0001   .0012   .0261

UNCT:   1181.8  6650.0  2017.9   233.9    21.5   373.8    20.1     8.6     3.4      .6     5.6   888.5

UNBG:     12.7    26.6    21.3     8.5     9.8     5.9     1.6     3.3     1.7     8.2     4.4    14.3

ZCOR:   1.0787  1.2848  1.3920  2.0257  1.1032  1.1792  1.1716  1.2417  1.4051  1.1759  1.1996  1.5501

KRAW:    .4210   .9999   .7588   .1354   .0131   .1130   .0162   .0032   .0043   .0002   .0017   .0551

PKBG:    93.74  250.75   95.75   28.67    3.21   64.45   13.24    3.68    3.09    1.08    2.35   63.00

INT%:      .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00

St  298 Set   2 Glass, VG-2 USNM 111240/52

TakeOff =  40   KiloVolts =  15    Beam Current =  20   Beam Size =  10

Juan de Fuca Basaltic Glass

Number of Data Lines:  5               Number of 'Good' Data Lines:  5

First/Last Date-Time: 11/11/2004 04:55:55 PM to 11/11/2004 05:06:58 PM

Average Total Oxygen:       43.790     Average Total Weight%:   99.343

Average Calculated Oxygen:  43.790     Average Atomic Number:   12.809

Average Excess Oxygen:        .000     Average Atomic Weight:   22.070

Average ZAF Iteration:        3.00     Average Quant Iterate:     3.00

Results in Elemental Weight Percents

SPEC:        O

TYPE:     CALC

AVER:  43.7900

SDEV:   .10338

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg

BGDS:      MAN     MAN     MAN     MAN     LIN     MAN     MAN     LIN     LIN     LIN     LIN     MAN

ABS%:    -3.48  -21.61  -26.61  -49.62     .00    -.29   -2.50     .00     .00     .00     .00  -35.87

TIME:    10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00   10.00

ELEM:       Ca      Si      Al      Na       K      Fe      Ti       S       P      Ni      Mn      Mg   SUM

15 7.65475 23.7958 7.39514 1.98209  .16231 9.12237  .96974  .16762  .08598  .00507  .04794 3.99781 99.1256

16 7.64931 23.7441 7.39264 1.96989  .17020 8.81813 1.13566  .14287  .09453  .02029  .13018 4.00919 98.9833

17 7.67023 23.6587 7.42488 1.93465  .15238 9.02186 1.09259  .13852  .08592  .00000  .21747 4.07332 99.1748

18 7.70080 23.7525 7.39467 2.10569  .15194 9.14024 1.02494  .14412  .09455  .04818  .14112 4.10122 99.6734

19 7.77206 23.7884 7.42423 2.08669  .17793 9.05015 1.03510  .13045  .12601  .01014  .20417 4.02589 99.7585

AVER:  7.68943 23.7479 7.40631 2.01580  .16295 9.03055 1.05161  .14472  .09740  .01674  .14818 4.04149 99.3431

SDEV:   .05034  .05459  .01668  .07572  .01130  .12850  .06409  .01388  .01656  .01911  .06775  .04409

SERR:   .02251  .02441  .00746  .03386  .00505  .05747  .02866  .00621  .00741  .00855  .03030  .01972

%RSD:       .7      .2      .2     3.8     6.9     1.4     6.1     9.6    17.0   114.2    45.7     1.1

PUBL:  7.94739 23.7501 7.44119 1.94365  .15773 9.20329 1.10907  .13500  .08728   n.a.   .17038 4.04631 99.7554

%VAR:    -3.25    -.01    -.47    3.71    3.31   -1.88   -5.18    7.20   11.59     .00  -13.03    -.12

DIFF:  -.25796 -.00219 -.03488  .07215  .00522 -.17274 -.05747  .00972  .01012  .00000 -.02220 -.00482

STDS:      358     298     374     336     374     395      22     757     285      28      25      12

STKF:    .1693   .1849   .0701   .0735   .1132   .6779   .5547   .3690   .1599   .7502   .7341   .4736

STCT:   2807.0  6650.6  2659.3  1728.1  1650.2  3309.5  1241.8  2716.6   778.3  3452.2  3325.8 16139.1

UNKF:    .0713   .1848   .0532   .0100   .0015   .0766   .0090   .0012   .0007   .0001   .0012   .0261

UNCT:   1181.8  6650.0  2017.9   233.9    21.5   373.8    20.1     8.6     3.4      .6     5.6   888.5

UNBG:     12.7    26.6    21.3     8.5     9.8     5.9     1.6     3.3     1.7     8.2     4.4    14.3

ZCOR:   1.0787  1.2848  1.3920  2.0257  1.1032  1.1792  1.1716  1.2418  1.4051  1.1759  1.1996  1.5501

KRAW:    .4210   .9999   .7588   .1354   .0131   .1130   .0162   .0032   .0043   .0002   .0017   .0551

PKBG:    93.78  250.62   95.70   28.66    3.21   64.45   13.27    3.68    3.09    1.08    2.35   63.00

INT%:      .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00     .00